ESP32 Multi-Sensor Integration Guide

YF-S201 Flow Meter  |  MH-Z19E CO2 Sensor  |  SEN0169 pH Sensor

With Wi-Fi AP Dashboard for Real-Time Monitoring

Rev 1.0  —  April 2026

# 1. System Overview

This document covers the wiring, firmware, and calibration for a three-sensor monitoring system built on an ESP32 DevKit. The ESP32 broadcasts its own Wi-Fi Access Point and serves a live dashboard accessible from any phone or laptop browser.

| Component | Model | Interface | Power |
| --- | --- | --- | --- |
| Flow Meter | YF-S201 | Digital Pulse | 5V (from VIN) |
| CO2 Sensor | MH-Z19E (NDIR) | UART (Serial2) | 5V (from VIN) |
| pH Sensor | DFRobot SEN0169 | Analog (ADC) | 3.3V |
| Microcontroller | ESP32 DevKit | — | USB 5V |

Important Note on pH Sensor Power: The SEN0169 signal conditioner board is rated for 5V operation. Running it at 3.3V compresses the output voltage range and alters the calibration slope from the factory default of 3.5 to approximately 5.3. Accuracy may be reduced from the rated ±0.1 pH. Calibration with buffer solutions is mandatory.

# 2. Wiring Connections

## 2.1  MH-Z19E CO2 Sensor

Uses ESP32 Hardware Serial2. No voltage divider needed — ESP32 is natively 3.3V logic and can read the sensor’s 3.3V TX output directly. The sensor’s internal IR lamp requires 5V power.

| Sensor Pin | Connect To | Notes |
| --- | --- | --- |
| Vin (Power) | ESP32 VIN (5V) | Must be 5V for IR lamp |
| GND | ESP32 GND | — |
| TX | ESP32 GPIO 16 | Hardware RX2 |
| RX | ESP32 GPIO 17 | Hardware TX2 |

## 2.2  YF-S201 Flow Meter

The YF-S201 is powered at 5V, so its signal output is 5V pulses. A 10kΩ / 20kΩ voltage divider is required on the yellow signal wire to step it down to ~3.3V before it reaches the ESP32 GPIO pin.

| Sensor Pin | Connect To | Notes |
| --- | --- | --- |
| Red (VCC) | ESP32 VIN (5V) | — |
| Black (GND) | ESP32 GND | — |
| Yellow (Signal) | ESP32 GPIO 27 | Through 10k/20k voltage divider |

### Voltage Divider Wiring (YF-S201 Signal)

Yellow wire → 10kΩ resistor → junction point → GPIO 27

Junction point → 20kΩ resistor → GND

Output = 5V × 20k/(10k+20k) = 3.33V

## 2.3  DFRobot SEN0169 pH Sensor

Powered at 3.3V. Analog signal connects directly to GPIO 32 — no voltage divider needed since the board output cannot exceed its own supply voltage.

| Sensor Pin | Connect To | Notes |
| --- | --- | --- |
| + (VCC) | ESP32 3.3V | Powered at 3.3V (see note above) |
| – (GND) | ESP32 GND | — |
| A (Analog Out) | ESP32 GPIO 32 | ADC1_CH4 — direct connection |

GPIO 32 is an ADC1 channel, which remains functional when Wi-Fi is active (ADC2 channels are blocked by Wi-Fi). The BNC connector on the board accepts the industrial pH electrode.

# 3. ESP32 Pin Summary

| GPIO | Function | Sensor | Signal Type |
| --- | --- | --- | --- |
| GPIO 16 | Serial2 RX | MH-Z19E TX | UART 9600 baud |
| GPIO 17 | Serial2 TX | MH-Z19E RX | UART 9600 baud |
| GPIO 27 | Interrupt Input | YF-S201 Signal | Digital pulse (3.3V via divider) |
| GPIO 32 | ADC1_CH4 | SEN0169 Analog | Analog 0–3.3V |
| VIN (5V) | Power Out | MH-Z19E + YF-S201 | 5V from USB |
| 3V3 | Power Out | SEN0169 | 3.3V regulated |
| GND | Common Ground | All sensors | — |

# 4. Wi-Fi Dashboard Access

- Power the ESP32 via USB.

- On your phone/laptop, connect to Wi-Fi network: Sensor_Dashboard

- Password: password123

- Open a browser and navigate to: 192.168.4.1

- The dashboard updates all three sensor readings every 1 second.

# 5. pH Sensor Calibration

Since the SEN0169 is powered at 3.3V instead of the rated 5V, factory defaults do not apply. You must calibrate using buffer solutions.

## 5.1  Electrode Conditioning

If the electrode is new or has not been used recently, soak the glass bulb and sand core in 3N KCl solution for at least 8 hours before calibration.

## 5.2  Two-Point Calibration Procedure

Step 1 — Midpoint (pH 7.00): Short the BNC connector or immerse the electrode in pH 7.00 buffer. Observe the voltage on the serial monitor or phone dashboard. It should read approximately 1.65V at 3.3V power. Note the pH reading. If it is not 7.00, calculate the difference and set PH_OFFSET in the code. For example, if it reads 7.15, set PH_OFFSET to –0.15.

Step 2 — Acid Point (pH 4.00): Rinse the electrode with deionized water, then immerse in pH 4.00 buffer. Wait one minute for stabilization. Adjust the gain potentiometer on the SEN0169 signal conditioner board until the dashboard reads 4.00. If the pot adjustment is insufficient, tweak PH_SLOPE in firmware.

Step 3 — Alkaline Verification (Optional, pH 9.18): For higher accuracy in alkaline ranges, rinse and immerse in pH 9.18 buffer. The reading should be close to 9.18 without further adjustment if the two-point calibration was done correctly.

## 5.3  Calibration Constants in Code

The two key defines in the firmware:

#define PH_OFFSET 0.00   // Adjust after pH 7.00 buffer test

#define PH_SLOPE  5.3    // ~3.5 * (5.0/3.3) for 3.3V operation

The slope of 5.3 is a starting estimate. Derive exact values from your two-point buffer measurements using: slope = (pH2 - pH1) / (V2 - V1) where V is the measured analog voltage.

# 6. Complete ESP32 Firmware

Copy this entire code into Arduino IDE. Board: ESP32 Dev Module. Upload speed: 115200. Serial monitor: 115200 baud.

#include <WiFi.h>

#include <WebServer.h>

// --- PIN DEFINITIONS ---

const int flowSensorPin = 27;

const int phSensorPin = 32;

// MH-Z19E: GPIO 16 (RX2), GPIO 17 (TX2) via Serial2

// --- WI-FI ACCESS POINT ---

const char* ssid = "Sensor_Dashboard";

const char* password = "password123";

WebServer server(80);

// --- GLOBAL SENSOR DATA ---

volatile int pulseCount = 0;

float currentFlowRate = 0.0;

int currentCO2 = 0;

float currentPH = 0.0;

float currentPHVoltage = 0.0;

unsigned long oldTime = 0;

// --- pH CALIBRATION (3.3V POWER) ---

#define PH_OFFSET 0.00

#define PH_SLOPE  5.3

#define PH_ARRAY_LENGTH 40

int pHArray[PH_ARRAY_LENGTH];

int pHArrayIndex = 0;

// --- FLOW SENSOR ISR ---

void IRAM_ATTR pulseCounter() {

pulseCount++;

}

// --- pH averaging (removes outliers) ---

double averageArray(int* arr, int number) {

int i, maxVal, minVal;

double avg;

long amount = 0;

if (number <= 0) return 0;

if (number < 5) {

for (i = 0; i < number; i++) amount += arr[i];

return (double)amount / number;

}

if (arr[0] < arr[1]) { minVal = arr[0]; maxVal = arr[1]; }

else { minVal = arr[1]; maxVal = arr[0]; }

for (i = 2; i < number; i++) {

if (arr[i] < minVal) { amount += minVal; minVal = arr[i]; }

else if (arr[i] > maxVal) { amount += maxVal; maxVal = arr[i]; }

else { amount += arr[i]; }

}

return (double)amount / (number - 2);

}

// --- MH-Z19E UART read ---

int readCO2() {

byte cmd[9] = {0xFF,0x01,0x86,0x00,0x00,0x00,0x00,0x00,0x79};

byte response[9];

while (Serial2.available() > 0) { Serial2.read(); }

Serial2.write(cmd, 9);

unsigned long timeout = millis();

while (Serial2.available() < 9) {

if (millis() - timeout > 500) return -1;

delay(5);

}

Serial2.readBytes(response, 9);

byte crc = 0;

for (int i = 1; i < 8; i++) { crc += response[i]; }

crc = 255 - crc; crc++;

if (crc == response[8]) {

return (response[2] * 256) + response[3];

}

return -1;

}

// --- DASHBOARD HTML (stored in flash) ---

const char index_html[] PROGMEM = R"rawliteral(

<!DOCTYPE HTML><html>

<head>

<meta name="viewport" content="width=device-width, initial-scale=1">

<title>AirForest Telemetry</title>

<style>

body { font-family: Arial, sans-serif; background: #121212;

color: #fff; text-align: center; margin: 0; padding: 20px; }

h2 { color: #bb86fc; font-size: 22px; margin-bottom: 20px; }

.card { background: #1e1e1e; border-radius: 15px; padding: 18px;

margin: 12px auto; width: 90%; max-width: 400px;

box-shadow: 0 4px 8px rgba(0,0,0,0.3); }

.value { font-size: 44px; font-weight: bold; color: #03dac6; margin: 8px 0; }

.unit { font-size: 16px; color: #aaa; }

.label { font-size: 14px; text-transform: uppercase;

letter-spacing: 1px; color: #888; }

.error { color: #cf6679; }

.voltage { font-size: 12px; color: #666; margin-top: 4px; }

</style>

</head>

<body>

<h2>AirForest Sensor Dashboard</h2>

<div class="card">

<div class="label">Air Flow Rate</div>

<div class="value" id="flowVal">--</div>

<div class="unit">Liters / Min</div>

</div>

<div class="card">

<div class="label">CO2 Concentration</div>

<div class="value" id="co2Val">--</div>

<div class="unit">PPM</div>

</div>

<div class="card">

<div class="label">pH Level</div>

<div class="value" id="phVal">--</div>

<div class="unit">pH</div>

<div class="voltage" id="phVolt">--</div>

</div>

<script>

setInterval(function() {

fetch('/data').then(r => r.json()).then(data => {

document.getElementById("flowVal").innerText = data.flow.toFixed(2);

if(data.co2 > 0) {

document.getElementById("co2Val").innerText = data.co2;

document.getElementById("co2Val").className = "value";

} else {

document.getElementById("co2Val").innerText = "WAIT";

document.getElementById("co2Val").className = "value error";

}

document.getElementById("phVal").innerText = data.ph.toFixed(2);

document.getElementById("phVolt").innerText =

"Analog V: " + data.phv.toFixed(3) + " V (raw: " + data.phr + ")";

});

}, 1000);

</script>

</body></html>)rawliteral";

void setup() {

Serial.begin(115200);

Serial2.begin(9600, SERIAL_8N1, 16, 17);

pinMode(flowSensorPin, INPUT_PULLUP);

attachInterrupt(digitalPinToInterrupt(flowSensorPin), pulseCounter, FALLING);

analogReadResolution(12);

analogSetAttenuation(ADC_11db);

WiFi.softAP(ssid, password);

Serial.print("Dashboard IP: ");

Serial.println(WiFi.softAPIP());

server.on("/", []() {

server.send(200, "text/html", index_html);

});

server.on("/data", []() {

String json = "{\"flow\":" + String(currentFlowRate) +

",\"co2\":" + String(currentCO2) +

",\"ph\":" + String(currentPH, 2) +

",\"phv\":" + String(currentPHVoltage, 3) +

",\"phr\":" + String((int)averageArray(pHArray, PH_ARRAY_LENGTH)) + "}";

server.send(200, "application/json", json);

});

server.begin();

Serial.println("System online. Connect to 'Sensor_Dashboard', open 192.168.4.1");

oldTime = millis();

}

void loop() {

server.handleClient();

// pH: continuous sampling every 20ms

static unsigned long phSampleTime = millis();

if (millis() - phSampleTime > 20) {

pHArray[pHArrayIndex++] = analogRead(phSensorPin);

if (pHArrayIndex == PH_ARRAY_LENGTH) pHArrayIndex = 0;

phSampleTime = millis();

}

// 1-second telemetry cycle

if ((millis() - oldTime) > 1000) {

// Flow sensor

detachInterrupt(digitalPinToInterrupt(flowSensorPin));

currentFlowRate = (float)pulseCount / 7.5;

pulseCount = 0;

attachInterrupt(digitalPinToInterrupt(flowSensorPin), pulseCounter, FALLING);

// CO2 sensor

currentCO2 = readCO2();

// pH sensor (3.3V powered, direct read)

float avgRaw = averageArray(pHArray, PH_ARRAY_LENGTH);

currentPHVoltage = avgRaw * 3.3 / 4095.0;

currentPH = PH_SLOPE * currentPHVoltage + PH_OFFSET;

// Serial output

Serial.print("Flow: "); Serial.print(currentFlowRate, 2);

Serial.print(" L/min | ");

if (currentCO2 > 0) {

Serial.print("CO2: "); Serial.print(currentCO2);

Serial.print(" ppm | ");

} else { Serial.print("CO2: waiting | "); }

Serial.print("pH: "); Serial.print(currentPH, 2);

Serial.print(" (V="); Serial.print(currentPHVoltage, 3);

Serial.print(", raw="); Serial.print((int)avgRaw);

Serial.println(")");

oldTime = millis();

}

}

# 7. Important Notes

## 7.1  MH-Z19E ABC (Auto Baseline Calibration)

The MH-Z19E ships with ABC enabled. Every 24 hours it assumes the lowest reading is fresh air (~400 ppm). For sealed environments like photobioreactors where CO2 never drops to ambient, disable ABC by sending this command once during setup:

byte abcOff[9] = {0xFF,0x01,0x79,0x00,0x00,0x00,0x00,0x00,0x86};

Serial2.write(abcOff, 9);

## 7.2  YF-S201 Limitation for Air Flow

The YF-S201 is a water flow meter. Air is ~800x less dense than water and may not generate enough force to spin the internal turbine reliably. Expect inaccurate or zero readings for low-velocity air. For precise air flow measurement, consider a hot-wire mass air flow sensor.

## 7.3  ESP32 ADC Limitations

The ESP32 ADC is 12-bit but is known to be non-linear at the extremes (below ~0.1V and above ~3.1V). For the pH sensor operating in the typical 5–10 pH range, the voltage stays in the linear midband and this is generally not an issue. For higher precision, consider an external ADS1115 16-bit ADC.

# 8. Document Revision History

| Rev | Date | Changes |
| --- | --- | --- |
| 1.0 | April 2026 | Initial release — ESP32 + YF-S201 + MH-Z19E + SEN0169 with WiFi AP dashboard |